## Video System
Indicative Specifications / Evaluation Form
Item No. | Medical Video Recorder and Integration System | Qty. | Yes/No | To Be Filled By Vendor | To Be Filled By Hospital
|  |  |  | Supporting Documents Reference | Met
1 | Technical Specs |  |  |  | Yes | No
Technical specs for the Medical Video Recorder and Integration System
1.1 | Medical Grade Video Recorder | 7
1.2 | Frame/design
1.2.0 | Full HD video recorder.
1.2.1 | Still Image capturing.
1.2.2 | 3D/2D recording.
1.2.3 | Two channel recording ( Ch1/Ch2). Capability to record from both at the same time.
1.2.4 | Large HDD supports long recording times Internal Hard Disc with min 4TB (4000GB) Capacity.
1.2.5 | Simultaneous Recording onto the internal hard disc and Two External Media (Blue-ray/DVD & USB or USB&USB & External Media Server).
1.2.6 | Built-in LCD displaying live signal and menu access.
1.2.7 | Compatible with DICOM.
1.2.8 | Remote Control and / or Foot Switch
1.2.9 | Touch panel to control all the features of the recorder and the simultaneous recording of the two channels.
1.2.10 | Network data transmission via CIFS with Streaming capability.
1.2.11 | Wide range of analog and digital video interfaces.
1.3 | Video Integration with Hospital Information Management System using HL7 | 1
1.3.0 | Content Management Solution is Central repository for video content shot in multiple ORs and procedure rooms for repurposing and reuse
1.3.1 | Highly efficient and secure content management allow storing and retrieving of all the video records.
1.3.2 | Content are easily searchable by hospital associated information system (HIS) and associated to the patient file.
1.3.3 | Can deliver a centralized imaging database even if connecting multiple O.R.s with varying camera system equipment.
1.3.4 | Can be expanded as the facility’s storage needs grow
1.3.5 | capability of live broadcasting with two way audio conferencing.
1.3.6 | Compatible IP TV (65” Ultra Full HD professional television) should be included in the proposal to access the recorded or live streaming procedures. The proposed wireless headset should be compatible to be used with the proposed television.
1.3.7 | Storage Retension as per UAE law for healthcare
1.3.8 | Storage should be sized to accomedate five years data with 20% growth every year
1.3.9 | Storage should support migration of data to the newer technology at the time of replacement/Upgrade
1.3.10 | The solution should have the ability to compress the data before it get stored
1.3.11 | Solution should have replication/backup of production data
1.3.14 | Required Hardware (servers & Storage) should be included in the proposals
1.3.15 | Solution should support the High availability
2 | Accessories
2.1 | Vendor to quote supply all mandatory accessories that will required for full intended functioning of the System.
3 | Consumables
3.1 | Vendor must include prices of all standard and optional consumables.
4 | Available Additional Options
No. | Options Description |  |  |  | REQUIRED OR NOT REQUIRED
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | A site visit is mandatory prior to submit the quotation, all preinstallation requirement should be discussed and included in the prices.
3 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) , IT team , for OT Nurses and Physicians (End User Training) is mandatory.
4 | Equipment Should Meets OR Exceeds IEC-60601-1-1, ISO, ANS, BS, JIS
5 | Equipment installed in the critical areas as OR, ER, ICU… should be medical grade and frequency interference compatible.
6 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
7 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
8 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
9 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering and IT Department.
10 | Comprehensive warrenty minimum 3 Years